FOX MODEL FT1 FLOW METER PROCUREMENT SPECIFICATION

- The flow meter shall operate on the Constant Delta Temperature (Constant Δ T) thermal mass principal.

- A DDC-Sensor™, or direct digitally controlled sensor, shall be standard.

- The Gas-SelectX® menu will offer pre-programmed and selectable gases and gas mixtures.

- The flow meter shall have a built-in display of flow rate, flow total, temperature, and elapsed time. The read-out shall utilize a backlit LCD display consisting of two lines each 16 characters.

- A 4-key keypad will be employed for user programming. Input parameters shall be protected by use of a password. Nonvolatile memory will retain the last totalizer value and user parameters.

- One 4-20mA output programmable for flow rate or temperature is required - HART communication option; a second output for pulse, RS485 Modbus RTU, or BACnet MS/TP option is selectable.

- The flow meter shall have a built-in microprocessor allowing field programmability of the 4mA setting, 20mA setting, pulse output setting, pipe diameter, zero flow cutoff, standard temperature and pressure (STP) and alarm settings.

- The flow meter shall have approvals from CE, FM/FMc, ATEX, and IECEx for use in potentially explosive atmospheres.

- The flow meter shall measure gas flows over a velocity range of 15-25,000 standard feet per minute. Sensor response time shall be 0.8 seconds for a one (1) time constant.

- In an operating temperature range of -40°F to 250°F, accuracy shall be ±1.0 percent of reading, ±0.2 percent of full scale for air and nitrogen calibrations; ±1.5 percent of reading, ±0.5 percent of full scale for other gases. Repeatability shall be ±0.2 percent of full scale.

- All wetted parts are to be 316SS utilizing an all welded design.

- All electronics to be mounted in a single NEMA 4X enclosure. Input power will be 12-24VDC.

- USB serial communication port is standard; the following communication options are also available: RS485 Modbus RTU or BACnet MS/TP.

- The manufacturer shall provide an NIST-traceable calibration certificate for the instrument.

- The flow meter will have the CAL-V™ calibration validation feature and internal self-diagnostics without requiring external equipment to evaluate meter performance.

- A CAL-V™ Calibration Validation Certificate can be printed upon completion of a CAL-V™ test initiated from the Fox FT1 View™ software.

- The instrument will be the Model FT1 manufactured by Fox Thermal, 399 Reservation Road, Marina, CA 93933 Phone: 831-384-4300, Email: sales@foxthermal.com, Website: www.foxthermal.com